## Sheet1
Input/Output | Connections | Arduino Mega Pin | Comments
Momentary Buttons (active LOW)
TV Change | Digital Input | 4 | Command change of tidal volume setting
Inspiration Time Change | Digital Input | 5 | Command change of inspiration time setting
Respiratory Rate Change | Digital Input | 6 | Command change of the respiratory rate setting
AC Mode Threshold Pressure Change | Digital Input | 7 | Command change of the pressure below the Low PEEP setting for which a breath will be given by the ventilator
Plateau Pause Time Change | Digital Input | 8 | Change hold time after inhalation
Future Version Button | Digital Input | 9 | No functions yet
High Pressure Alarm Setting | Digital Input | 22 | Command change to high pressure alarm set point
Low Pressure Alarm Setting | Digital Input | 24 | Command change to low pressue alarm set point
High PEEP Alarm Setting | Digital Input | 26 | Command change to high PEEP alarm set point
Low PEEP Alarm Setting | Digital Input | 28 | Command change to low PEEP alarm set point
Low PlateauPressure Setting | Digital Input | 30 | Command change to the low plateau pressure setting - used for disconnect alarm
Future Version 2 Button | Digital Input | 32 | No functions yet
Future Version 3 Button | Digital Input | 34 | No functions yet
Future Version 4 Button | Digital Input | 36 | No functions yet
Limit Switch | Interrupt | 10
Alarm Silence Button | Interrupt | 18
Rotary Encoder
Encoder Channel A | Interrupt 3 | 2
Encoder Channel B | Interrupt 4 | 3
Encoder Momentary Button | Interrupt 5 | 19 | Confirms change to setting
Toggle Switches
AC/VC Mode Switch | Digital Input 13 | 38 | Selects between AC and VC modes
LCD Screens
Parameter LCD Screen | Param RS Pin | 53
| Param Enable Pin | 51
| Param DB4 | 47
| Param DB5 | 45
| Param DB6 | 43
| Param DB7 | 41
Alarm LCD Screen | Alarm RS Pin | 53
| Alarm Enable Pin | 49
| Alarm DB4 | 47
| Alarm DB5 | 45
| Alarm DB6 | 43
| Alarm DB7 | 41
Sensors
Pressure Transducer | SDA | 20 (SDA)
| SCL | 21 (SCL)
Motro Controller
RoboClaw | RX | 17 (RX2)
| TX | 16 (TX2)
Outputs
Alarm Buzzer | Digital Output | 11
Alarm LCD | Digital Output | 12
Alarm Relay | Digital Output | 13