## Alarm Logic
What Alarm | When Does it Occur? | What Action Does the Software Take? | How Long does the Alarm Last? | Once Dismissed/Silenced
High Pressure Alarm | Measured pressure is above the set high pressure limit during inhalation | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
Low Pressure Alarm | Peak inspiratory pressure is below the set low pressure limit during inhalation | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
High PEEP Alarm | Measure PEEP pressure is above the set high PEEP pressure limit | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
Low PEEP Alarm | Measure PEEP pressure is below the set low PEEP pressure limit | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
Disconnect Alarm | Low Pressure Alarm is triggered | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
High Temperature Alarm | Motor controller temperature is above the set high temperature limit | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
Apnea Alarm | AC inhale wait time is above the AC threshold time | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
Device Failure Alarm | Motor position no responding to commanded position OR Watchdog timeout | Stop all motion; demand restart | Indefinite | Demand device is restarted
Variable in red = measured by sensor or controller each loop
Variable in purple = set by user via the UI
Varible in blue = internal or hard-coded value