## Alarm Logic
What Alarm | When Does it Occur? | What Action Does the Software Take? | How Long does the Alarm Last? | If Dismissed/Silenced by User
High Pressure Alarm | Measured pressure is above the set high pressure limit** during inhalation | Audible/visual alarm; stop inhale; jump to exhale; reset to try inhale again | Until next breath | Disable for 1-minute
Low Pressure Alarm | Peak inspiratory pressure is below the set low pressure limit during inhalation | Audible/visual alarm; continue ventilating | Until next breath | Disable for 1-minute
High PEEP Alarm | Measure PEEP pressure is above the set high PEEP pressure limit | Audible/visual alarm; continue ventilating | Until next breath | Disable for 1-minute
Low PEEP Alarm | Measure PEEP pressure is below the set low PEEP pressure limit | Audible/visual alarm; continue ventilating | Until next breath | Disable for 1-minute
Disconnect Alarm | Plateau pressure is below the set low plateau pressure limit | Audible/visual alarm; continue ventilating | Until next breath | Disable for 1-minute
High Temperature Alarm | Motor controller temperature is above the set high temperature limit | Audible/visual alarm; continue ventilating | Until next breath | Disable for 1-minute
Apnea Alarm | AC inhale wait time is above the AC threshold time | Audible/visual alarm; continue ventilating | Until dismissed | Disable for 1-minute
Mechanism Failure Alarm | Motor position not responding to commanded position | Stop all motion; demand restart | Indefinite | Demand device is restarted
Variable in red = measured by sensor or controller each loop
Variable in purple = set by user via the UI
Varible in blue = internal or hard-coded value
** Note that the high pressure limit should be hard coded to a maximum of 40 cm H2O